Prerequisites

You are familiar with NIS-elements – use help tool within software and in the manual

Note this script is based on NIS-elements template JOBS scripts

Use the NIS-Elements help tool to work out how and why you would want to use JOBS scripts.

Contents

Introduction	1

Run acquisition JOBS script	2

defined dOPM view 1&2 refocus offsets	2

Set z_retract parameter	2

Select multi-well plate	3

Set timelapse parameters	4

Load position list	4

Define dOPM volume acquisition – spectral channels	5

Define dOPM volume acquisition - z-scan settings	5

Wizard completed	6

Run acquisition	6

# Introduction

NOTE: this script makes use of a logical device called z-piezo in NIS-elements but not for its intended purpose. Normally this device is attached to the microscope objective for fast z-positioning. Here we use it for remote refocussing within dOPM's external optics on the left-hand port of the microscope frame. With this in mind we only use z-piezo when doing dOPM volumes scans.

- Any other operations involving z-positions i.e., the normal use of the microscope, with the right-hand port should be carried out with z-piezo set to the home position which is zero.

- For example, if we want to record a position list of interesting locations in the sample, we first make sure the z-piezo is set to zero otherwise NIS-elements sets the target position as main z-drive value + z-piezo drive value i.e. ZTotal = Z1 + Z2. This is not physically meaningful since the z-piezo drive (Z2) is not attached to the microscope objective. For any position list we always work with ZTotal = Z1

Run acquisition JOBS script

- Run the ‘dOPM_acquisition’ JOBS script

- Use the NIS-Elements help tool to work out how and why you would want to:

  - Use JOBS explorer

defined dOPM view 1&2 refocus offsets

- Enter the zero remote refocus offset values for dOPM

# Set z_retract parameter

- Set the z_retract variable to a number that avoids the chance of the objective colliding with the plate during stage movements.

- Here a conservative 300µm is used but depends on the plate being used – see JOBS script ‘GetPlateZProfile’ for a way to measure plate flatness.

# Select multi-well plate

- choose correct multi-well plate definition

Set timelapse parameters

- choose whether time-lapse or not and the frequency of time points

- need to choose frequency that is longer than acquisition time for position list

- Use the NIS-Elements help tool to understand timing requirements in timelapse imaging

# Load position list

- load or define position list xml or csv file

Define dOPM volume acquisition – spectral channels

- Setup the dOPM ND acquisition parameters - lambda tab

  - suitable dOPM optical configurations are selected

  - advanced tab at bottom that following command boxes are ticked and filled in with

  - execute command before capture StgMovePiezoZ(dOPM_offset,0)

  - execute command after capture StgMovePiezoZ(dOPM_offset,0)

# Define dOPM volume acquisition - z-scan settings

- setup the dOPM ND acquisition parameters - z tab

  - NOTE: there are two tabs to use and the others should be empty and unchecked

  - for the z tab, set z scanning device to either

    - NIDAQ piezo

      - NIDAQ piezo is software timed and and there will be no motion blur but relatively slow.

    - Triggered NIDAQ piezo

      - Triggered NIDAQ piezo is hardware timed and there will be motion blur but relatively fast.

  - NOTE: stick with option 1 above if unsure. Choice depends on height of image in pixels, exposure time, z step size

# Wizard completed

- once all steps in the JOB script setup have been addressed the green button is enabled and the acquisition can be started

# Run acquisition

- Acquisition is initiated by pressing green button. Progress bars at the top indicate acquisition timing details e.g. above

  - Timelapse – x10 time points

  - Points2 –  x8 xyz positions

  - Repeat1 – View 1&2 dOPM

  - Zstack – View1&2 dOPM zstack planes

Assumptions

- Acquiring data with both dOPM views

- Well locations used are compatible with 60x objective and any water immersion cap on top – recall that 60x objective cannot typically reach the edges of a plate. Do not proceed with any 60x based experiment if this is not clear to user and ask for help.